## Slide 1
What we built
A standalone tPBM device, the software around it, and a way of working that kept it safe.
The device
Arduino-driven 850 nm LED array, pulsed at 10 Hz or 40 Hz
Heater for the warmth-only control condition
Skin temperature sensor with a 40 °C automatic cut-off
The software
Browser task page runs the session and the reaction-time test
Every event is sent as a marker into the EEG recording
Firmware logs each light pulse and temperature reading
How we worked
Every change tested on the real board before use
Bench scripts with automatic abort limits
Calibration stored as files, not typed by hand
Wire & check
Flash
Test on hardware
Calibrate
Run session
Notes: Three things: the physical device, the software that runs a session and tags the EEG, and the working loop. The loop at the bottom is the point: nothing goes on a participant until the physical board has confirmed it.

## Slide 2
Where we are
The device works on the bench. A few things remain before the first participant.
10 / 40 Hz
Pulse rate measured exactly, no dropped pulses
0.26 °C
Overshoot of the warmth control (was 2.75 °C)
40 °C
Hard cut-off, verified: heater stops on command
Solved along the way
LED strip turned out to be 24 V: power stage redesigned
Heater runaway traced to wiring, not firmware, and fixed
EEG markers were being dropped: now sent as numeric codes
Next steps
Diagnose LED strip heating seen in long bench runs
Add a hardware thermal fuse and measure the light dose
Pilot session with EEG, then start participants
Notes: Top row: the three numbers that matter, all measured on the real board. Bottom left: the three problems worth telling as stories. Bottom right: what stands between here and the first participant.